BM15D Digital Multimeter Operating Instruction

Ⅰ.General introduction

Welcome to use this product!

This product is one kind of 3 1/2 portable multipurpose automatic measuring range instrument, may measure the AC/ DC voltage, the AC/ DC current, resistance, frequency, capacitance, temperature, continuity test,  diode. The measuring appliance is loaded with guards against the surge electric discharge, protect the appliance to be more effective. This measuring appliance structure is exquisite, the operation is easy, it is your ideal test service tool!

Ⅱ.Safety Rules and Notes

The instrument design conforms to EN1010-1:2010 600V CATIII,

EN61326-1:2013 Safety requirements for EMC standards. Please read this manual carefully before use.

Warning, be careful!

Danger of being hit by high-pressured electric!

Dual insulation protection.

2.2 when measure, do not surpass the greatest stipulated input value.

2.2 do not surpass 10V voltage to the input end, except the voltage grade

2.4 In the process of measuring, do not turn switch to change the measuring range at random, in case to destroy the measuring appliance.

2.5The measuring appliance can display the mark while the voltage is bigger than DC60V and AC30V, remind the user that the measured voltage has surpassed the safety voltage, please operate carefully.

2.6measuring appliances should avoid the straight sunlight, the high temperature, and moisture.

2.7after use, must release the power switch to turn off the power

2.8 if it doesn’t use for a long time, should take out the battery, in case the battery leaks to damage the parts.

Ⅲ. Features

3.1General Features

3.1.1 take the CMOS big scale integrated circuit as the core, in AC / DC voltage, the AC/DC electric current, the resistance, the frequency and the electric capacity measure it can automatically transform the measuring range, making it more convenient.

3.1.2 greatest display: 1999

3.1.3 automatic cathode display: Displays " - "

3.1.4 batteries insufficient display: Displays “ ”.

3.1.5 Auto power OFF

After turning on the instrument and without turning the function switch or pressing any button, the instrument will automatically enter into sleep mode after 15 minutes,to save battery power.when it is in the sleep mode you can press the any button to wake up the instrument. If you don't need the automatic sleep mode, you should hold down the FUNCE button to turn on the instrument, and then  the symbol"" will not be display on the LCD.

3.1.7 working condition：0C~40C，75%RH（max）

3.1.8 storage environment：-10C~60C，80%RH（max）

3.1.9 battery：AAA1.5V×2

3.1.10Fuse: F 200mA, 10A (size Φ3.8×10mm).

3.1.11 external dimensions ：152(L）73(W）48（H）mm

3.1.12 weight：approximately 227g（contain battery）

3．2Technical index

Accuracy: ± (% reading + word count) , Guarantee period of one year

Guaranteed accuracy Temperature: 23 ° C ± 5 ° C

Relative humidity: < 70%.

3．2．1 DCV

| Range | Accuracy | Resolution |
| --- | --- | --- |
| 200mV | ±(0.5% of rdg + 5 digit ) | 0.1mV |
| 2V |  | 1mV |
| 20V |  | 10mV |
| 200V |  | 100mV |
| 600V |  | 1V |

Input resistance：about 10MΩ.

Overload protection :DC/AC  600V.

3．2．2 ACV

| Range | Accuracy | Resolution |
| --- | --- | --- |
| 2V | ±(1.2% of rdg + 5 digit ) | 1mV |
| 20V |  | 10mV |
| 200V |  | 100mV |
| 600V |  | 1V |

Input resistance ：about10MΩ。 Frequency：10Hz~1kHz.

display：TRUE RMS(sinusoidal waveform RMS calibration).

Overload protection: DC/AC 600V.

3.2.3 DCA

| range | Accuracy | Resolution |
| --- | --- | --- |
| 200μA | ±(1% of rdg + 5 digit ) | 0.1μA |
| 2000μA |  | 1μA |
| 20mA | ±(1.5% of rdg + 5 digit ) | 0.01mA |
| 600mA |  | 0.1mA |
| 2A | ±(2% of rdg + 5 digit ) | 1mA |
| 10A |  | 10mA |

Overload protection ：μA/mA:F200mA fuse ，10A： F10 A fuse。

greatest input current：10A（less than 10 seconds）。

voltage drop measure：full measure range is 200mV。

3.2.4 ACA

| range | Accuracy | Resolution |
| --- | --- | --- |
| 200μA | ±(1% of rdg + 5 digit ) | 0.1μA |
| 2000μA |  | 1μA |
| 20mA | ±(1.8% of rdg + 5 digit ) | 0.01mA |
| 200mA |  | 0.1mA |
| 2A | ±(3% of rdg + 5 digit ) | 1mA |
| 10A |  | 10mA |

Overload protection：μA/mA:F200m A fuse，10A :F10A fuse。

Voltage drop measure：full measure range is 200mV(10A is 100mV)。

Frequency：10Hz~1kHz (Warning：Frequency for square wave accuracy is specified from 10Hz to 400Hz)。

display：TRUE RMS(sinusoidal waveform RMS calibration)。

greatest input electric current：10A（less than 10 seconds）。

3.2.5 resistance  

| range | Accuracy | Resolution |
| --- | --- | --- |
| 200Ω | ±(0.8% of rdg + 5 digit ) | 0．1Ω |
| 2KΩ |  | 1Ω |
| 20KΩ |  | 10Ω |
| 200KΩ |  | 100Ω |
| 2MΩ |  | 1KΩ |
| 20MΩ | ±(2% of rdg + 5 digit ) | 10KΩ |

Overload protection ：250Vvirtual value.

Open voltage approximately: 1V

3.2.6 CAP

| Range | Accuracy | Resolution |
| --- | --- | --- |
| 20nF | ±(3% of rdg + 5 digit ) | 0.01nF |
| 200nF |  | 0.1nF |
| 2uF |  | 1 nF |
| 20uF |  | 10nF |
| 200uF |  | 100nF |
| 2mF | ±(5% of rdg + 15 digit ) | 1uF |

Overload protection ：250Vvirtual value.

3.2.7 FREQ

| Range | Accuracy | Resolution |
| --- | --- | --- |
| 200Hz | ±(0.5% of rdg + 3digit ) | 0.1Hz |
| 2kHz |  | 1 Hz |
| 20kHz |  | 10 Hz |
| 200kHz |  | 10 0Hz |
| 2MHz |  | 1k Hz |
| 20MHz |  | 10k Hz |

Overload protection ：250Vvirtual value, input delicacy:2V。

Caution：if the measured frequency is above 30V, please press “FUNCE” key at AC electric voltage measuring range to get to the frequency function, then carry on measure。

3．2．8 Diode positive voltage

Display of similar diode positive voltage. Measuring condition：positive DC electric current 1mA，reverse DC voltage approximate2.2V。

3．2．9 Continuity Test

When the transited resistance is smaller than about 50,the buzzer beeps. Test condition:Open-circuit voltage is about 2V.

Ⅳ．Application method

4.1.1 FUNCE button

The FUNCE button: function choice button, in the trigger movement way. Cycle through the various functions on the Gear position.

4.1.2 RANGE button

RANGE button for the automatic / manual measuring range button , in the trigger movement way, before start the device it is at the automatic measuring range. It changes to manual measuring range when pish it. At manual measuring range it goes upward one gear with one touch, after it reaches the top gear, it goes downward to the lower-gear gradually with continuous press, and it takes turns. If push this button over 2 seconds, it changes back to the automatic measuring range condition.

4.1.3Data hold - backlight and flashlight light

4.1.3.1 DH

Data hold are maintained  In the trigger mode, when the button is

pushed, the displayed value is Holding ,the word “DH” on the

unchanged display; when push this button, the lock is the state is

released and enters the normal measurement state.

4.1.3.2 Backlight control and flashlight system

Push the DH button for more than 2 seconds to turn on the backlight and flashlight control signals. After the backlight signal is turned on, ppush the button for > 2 seconds to turn off the light. After the light is turned on,If you do not  ppush DH/LIGHT>2, the backlight will automatically turn off after 10 seconds.

4.2 DC/AC voltage measure

Turn the Rotary switch to “V”. Then plug black lead in “COM” socket, and plug red lead in “V/” socket. connect the test lead with the two ends of the circuit and then directly read the reading on the LCD display .

If manually choose measuring scope is needed, push RANGE key to choose.

Don't measure the peak voltage more than 600V, otherwise it might damage the instrument ,if the screen only displays OL, it means that the tested voltage is higher than 610V.

3. Push the FUNCE key in the ACV function to enter the voltage frequency measurement mode, it can measure the the frequency range from 10Hz to 20KHz with voltage higher than AC 2V~600V.

4.3 Measurement of DC/AC current

turn the Range switch to the range of electric current, this time the measuring appliance is set as the direct current automatic shift gears in advance. Insert the black lead into COM jack, and the red lead into 10A or μAmA  jack.

If one wants to measure DC, one can concatenate the table pen with the measured electric circuit, which can directly read the number on the liquid crystal display monitor; If one wants to measure alternating current, press FUNCE key to alternating current measuring range, then concatenate the table pen with the measured electric circuit to read the showed number.

If one needs to choose measuring scope manually, push RANGE.

If one does not know the measured electric current scope before survey, one should set the range key to the highest measuring range and adjust downward by the files.

When the display monitor only displays OL, that is to say, the measured electric current has surpassed the measuring range and the switch of measuring range needs to move the high one grade.

μA/mA jack represents that the maximum input current is 200mA. The overload inputs can burn the internal installation fuse out and should be replaced immediately.

10A jack, without a fuse, the time of measuring should be less than 10 seconds to avoid the heating of lines, which would influence accuracy.

4.4 Resistances

- turn the Range switch to the range to Ω position, and this time the measuring appliance will be set for the resistance measuring range.

- Insert red lead into the VΩ  jack, black lead into " COM " jack.

- Connect the pen with the beginnings and ends of the test circuit to read the resistance value.

- If the manual choice measuring range scope is needed, press RANGE to choose.

When the leads is overloaded input, display monitor can display " OL ".

4．5 Measurement of forward voltage of diode

（1）turn the  Range switch to “Ω// ”position. press “FUNCE” to  measuring range

（2）insert the red lead to “VΩ” jack，insert the black lead to “COM” jack.（red pen “+”）

（3）connect the test leads to the two ends of the measured to read positive voltage.

(4) when the diode is reverse connected or the input end leads the way, the display monitor can display " OL ".

4．6 Continuity  Testing

（1））Turn the  Range switch to “Ω// ”position, press “FUNCE” key to measuring range.

（2）Insert red lead to “VΩ” jack，insert black lead to “COM” Jack.

(3)If the measured resistance is less than about 50，the buzzer will beep. This is continuity testing..

Caution：

a. when the input end leads the way, it displays “OL”。

b. the measured circuit should be measured under the off-power condition ，because any overload signal would make the buzzer beep, thus cause wrong measure.

4. 7  Capacitance  measure

Warning！When measure the electric capacity, must guarantee the measured capacitor has sent the electricity out, if the big electric capacity contains the oversized non- electric capacity ingredient, possibly affects the measuring accuracy.

（1）turn the  Range switch to “”position .

ouchonly the figure’（2）insert the red lead to “VΩ” jack，insert black lead to “COM” jack.

（3）connect the test leads to the two ends of the measured capacitor，could get the capacity value.

（4）capacity position can not set the measuring range manually, no simulation strip display. when the capacity value is big, the measure may need about 10 seconds .

4．8 frequency

（1））turn the  Range switch to Hz measuring range.

（2）insert red lead to “VΩ” jack，insert black lead to “COM” jack.

（3）connect the test lead to the measured circuit ，get the frequency value. press “FUNCE” key，it displays “Hz”.

Caution: (1)Frequency grade can not set measuring range manually, no simulation strip display.

(2)If the measured frequency is above 30V, please press “FUNCE” key at “ACV” measuring range to frequency function for measure. By this method, it can bear 600V voltage，in case the voltage is too high to damage the meter .

4. 9 temperature measure

Turn the range switch to ℃/oF，insert the k type black lead to “COM” jack，insert red  lead  to “V” jack，the meter displays the temperature now . measure Fahrenheit, touch the “FUNCE” button to switch the “°F” function and measuring again.

Caution ：

When the meter without inserted with temperature tool, it displayd environment temperature（normal temperature）.Attached K type WRNM-010,the temperature is 250℃(482 oF)。

4.10 NON CONTACT VOLTAGE TESTING

Turn the Rotary switch to the “NCV” function, then the LCD display “EF”, leads without being used, place the upper NCV sensing probe to r the electric line, switch,or socket,the phase metal terminal ,When the AC voltage is detected, the meter displays "---". When the sensed voltage is higher, the number of "-" is displayed, the denser the sound accompanying the buzzer sounds.

Caution:

1.even if there is no indication, voltage may still exist. Do not judge the wire whether threre is voltage absolutely throught the non contact voltage testing, the testing may be effected by many factors such as the socket design,the insulation thickness and types etc.

2.Interference source of external environment,such as flash,motor etc,may false triggle the non contect voltage testing.

Ⅴ．Maintenance

warning！before open the cover or the battery cover, cut off the power source and test pen and any input signal, in case electric shock.

5．1 when the meter displays “”，must replace the battery. Open the battery cover, replace with the same type new battery to keep it work well.

5．2 keep the meter and test pen clean, dry and not damaged, could use the clean cloth or cleanser to clean the cover，do not use abrasive or solvent.

5．3 avoid damage, shake, shock，avoid high temperature and strong magnetic field.

5．4 should be corrected at least once per year

Ⅵ.	Accessories

6.1	Test lead: 1 set

6.2	users manual: 1 piece

6.3   cloth bag: 1piece

6.4   Temperature Sensor:1 piece